- PURPOSE

This procedure describes the apparatus, set up and measurement regime for the determination of the coefficient of retroreflected luminance of secondary reference plates for low angle hand-held retroreflectometers. The procedures described require a set of “primary” reference plates previously calibrated in accordance with TP991.

- REFERENCES

Operating manual for retroreflectometer

CIE No 54, Retroreflection Definition and Measurement

- DEFINITIONS

Retroreflection - Reflection in which the reflected rays are preferentially returned in directions close to the opposite of the direction of the incident rays, this property being maintained over wide variation of the direction of the incident rays.

Retroreflectivity (mcd /m² /lx) - The ratio of the luminance of a surface to the illuminance at the retroreflector on a plane perpendicular to the direction of the incident light. (Coefficient of retroreflected luminance)

Entrance Angle  - Angle characterising the angular position of the retroreflector with respect to the direction of the incident light.

Co-entrance Angle  - The complement of the entrance angle ° = (90-°)

Observation Angle  - Angle by which the direction of observation of the retrorelflector departs from the direction of the incident light. The angle between the illumination axis and the observation axis (CIE No 54, clause 2.4.7)

Primary Reference Plates – Flat glass plates manufactured to produce a range of retroreflectivity values, to provide a link to the primary references of length and luminous intensity.

Measurement Geometry – Entrance and observation angles defined by the European Committee for Standardisation (CEN), the US Federal Highway Administration (FHWA), and ASTM’s E- 1710 standard.

Low Angle Geometry – Entrance angle 88.76°, observation angle 1.05°.

- APPARATUS AND MATERIALS

4.1	Apparatus

A Dark Tunnel - with an effective length of at least 10m and stray light luminance of less than 1x10-3 cd m²

Low Angle Reflectometer – Entrance angle 88.76°, observation angle 1.05°. (MX 30)

Primary Plates Measurement Platform (DPTI – ID: 248-001) to provide accurate and repeatable location of the transfer retoreflectometer and the primary reference plates.

Primary Reference Plates – reference plates of varying retroreflectivity, calibrated in accordance with MAT-TP991, at the specified geometry.

Secondary Reference Plates Measurement Jig (DPTI – ID: 248-002) – to provide accurate and repeatable location of the retroreflectometers and secondary laboratory reference plates.

Secondary Reference Plates (DPTI – ID: LAB 30) a set of five reference plates of varying retroreflectivity, to accurately calibrate across from corresponding primary reference plates using the transfer retroreflectometer. Plates also used to establish data correction percentages for client reflectometers.

Secondary Client Reference Plates – MX30 timber boxed set of four reference plates, comprising ”zero”, “low”, “medium” and “high” retroreflectivity, issued with clients retroreflectometer.

Secondary Client Reference Plates – MX30 / Stripemaster metal boxed set of three reference plates, comprising “low”, “medium” and “high” retroreflectivity, issued with clients retroreflectometer.

Secondary Client Reference Plates – Zehntner set of three reference plates issued only, comprising “low”, “medium” and “high” retroreflectivity,

  - Materials

Nil.

- SAMPLING AND SAMPLE PREPARATION

5.1	Sampling

Nil.

  - Sample Preparation

Nil.

- PROCEDURE TO BE FOLLOWED

6.1	Calibration of Laboratory Secondary Reference Plates (LAB 30)

    - Clean primary reference plates in accordance with procedure outlined in appendix 5.

    - Prepare the retroreflectometer (*dedicated transfer instrument ) in accordance with the manufacturer’s operating instructions. Ensure that the stabiliser frame is securely fitted and that the light path dimensions are within specified requirements.

    - Calibrate instrument using manufacturer’s calibration factor.

    - Remove stabiliser frame.

    - Select a primary reference plate with a nominal retroreflectivity value of 0 -20 mcd/m2/lx at the required geometry, and record the plate identification number. Position plate and instrument on the primary plate measurement platform and take a reading. Adjust calibration factor until instrument reads the specified primary retroreflectivity reference value ±1 unit. Record calibration factor.

Note: Pre established calibration factors for the whole range of primary plates can be found in appendix 3. (Factors are for exclusive use with *dedicated transfer instrument MX30 1621/41 only, valid until November 2009.

    - Retain calibration factor, place instrument and corresponding “zero” laboratory plate into secondary measurement jig, measure and record the plate value.

Note: The range of plates has been measured; values are shown in appendix 2 (valid until June 2008).

    - Repeat steps 6.1.5 and 6.1.6 for the remaining 4 laboratory reference plates.

6.2	Calibration of Client MX 30 Instruments using Laboratory Secondary Reference Plates

    - Clean secondary reference plates in accordance with procedure outlined in appendix 5.

    - Prepare the retroreflectometer (client instrument) in accordance with the manufacturer’s operating instructions. Ensure that the stabiliser frame is securely fitted and that the light path dimensions are within specified requirements.

    - Calibrate instrument using manufacturer’s calibration factor.

    - Remove stabiliser frame.

    - Place retroreflectometer (client instrument) and “zero” laboratory secondary reference plate into secondary measurement jig. Ensure that the light patch covers the same area as the transfer instrument before measuring plate.

    - Measure “zero” plate and record value using worksheet - appendix 2, column “F”.

    - Compare measured retroreflectivity figure in column “F” to actual value in column “D”. Establish correction percentage by referring to correction chart – appendix 1, then record required ±% correction value using worksheet – appendix 2, column “G”.

    - Repeat steps 6.2.5 to 6.2.7 for the remaining 4 laboratory reference plates.

  - Measurement of Client Secondary Reference Plates (MX 30 Instruments with timber box, four plate sets)

    - Prepare secondary reference panels as described in appendix 5.

    - Prepare the retroreflectometer (client instrument) in accordance with the manufacturer’s operating instructions. Ensure that the stabiliser frame is securely fitted and that the light path dimensions are within specified requirements.

    - Calibrate instrument using manufacturer’s calibration factor.

    - Record client secondary reference plates ID. using worksheet – appendix 2, column “H”.

    - Place client reflectometer and “zero” secondary  plate into  timber box measuring platform, make sure that stabiliser frame feet do not interfere with instrument geometry, measure and record retroreflectivity value using worksheet – appendix 2, column “J”.

    - Repeat step 6.3.5 for the remaining 3 plates.

    - Identify and compare plates that are showing retroreflectivity values within 10% in columns “F” and “J”, 	then transpose corresponding correction percentage (refer to correction chart – appendix 1) ± percentage values from column ”G” to column ”K”.

    - Actual retroreflectivity values for client plates may be calculated (column “L”) from measured client plate values and ± correction percentages transposed from column’s ”G” into column “K”.

    - Copy data into MX30 report certificate, (TP993) appended.

  - Measurement of Client Secondary Reference Plates (MX 30 and Stripemaster Instruments with metal box, three plate sets)

    - Prepare secondary reference panels as described in appendix 5.

    - Prepare the retroreflectometer (client instrument) in accordance with the manufacturer’s operating instructions. Ensure that the stabiliser frame is securely fitted and that the light path dimensions are within specified requirements.

    - Calibrate instrument using manufacturer’s calibration factor.

    - Record client secondary reference plates ID. using worksheet – appendix 2, column “H”.

    - Place client reflectometer and “low” secondary plate into client metal box measuring platform, measure and record retroreflectivity value using worksheet – appendix 2, column “J”.

    - Repeat step 6.4.5 for the remaining 2 plates.

    - Identify and compare plates that are showing retroreflectivity values within 10% in columns “F” and “J”, 	then transpose corresponding correction percentage (refer to correction chart – appendix 1) ± percentage values from column ”G” to column ”K”.

    - Actual retroreflectivity values for client plates may be calculated (column “L”) from measured client plate values and ± correction percentages transposed from column’s ”G” into column “K”.

    - Copy data into MX30 / Stripemaster report certificate, (TP993) appended.

  - Measurement of Client Secondary Reference Plates (Zentner ZRM 6013 Instruments with own metal box, 3 plate sets)

    - Prepare secondary reference panels as described in appendix 5.

    - Prepare the retroreflectometer (dedicated transfer instrument) in accordance with the manufacturer’s operating instructions. Ensure that the stabiliser frame is securely fitted and that the light path dimensions are within specified requirements.

    - Remove stabiliser frame.

    - Record client secondary reference plates ID. Using worksheet – appendix 4, column “H”.

    - Pre-measure secondary client plates to establish approximate retroreflectivity values in order to select matching primary plate calibration factors.

    - Calibrate instrument using established calibration factor for corresponding retroreflectivity range to be measured. Refer to appendix 3 for calibration factors, enter factors required using worksheet – appendix 4, column “E”.

    - Place (dedicated transfer instrument) 1621/41 reflectometer and “low” secondary  plate into measuring platform DPTI 242 - 002, measure and record retroreflectivity value using worksheet – appendix 4, column “L”.

    - Repeat step 6.4.5 for medium and high plates.

    - Copy data into “Zehntner” report certificate, (TP993) appended.

7.0	PRECISION

The uncertainty of measurement has been determined as ± 7.62% for metal box, and ± 7.71% for timber box, sets of the test results at a confidence limit of 95% with a nominal coverage factor of 2, when calculated in accordance with the NATA accreditation requirements as detailed in ISO/IEC 17025.

Uncertainty of Measurement determined by and for DPTI Technical Services Group, Field Services Section.

8.0	DOCUMENTATION

8.1	Reports

Record results on form TP993.

Calibration results on forms TP993A – 1, 2, 3 & 4

MX 30 / Stripemaster Report No.  	Page 1 of 1

Laboratory Secondary Reference Set Identification Number:

Instrument Type:

Company:

Address:

Date Tested:

Tested By:

Calibration Ref.:

Test Data:

| Panel ID: | Actual retroreflectivity (mcd/m2/lx)* | Retroreflectivity Reading at Factory Calibration factor (mcd/m2/lx)* | Correction Required to Readings (± %) |
| --- | --- | --- | --- |
| H | L | J | K |

* Uncertainty of measurement has been determined as + 7.7% of the test result, at a confidence limit of 95% with a nominal coverage factor of 2

Uncertainty of Measurement determined by and for DPTI Technical Services Group, Field Services Section.

Test procedure TP993 recommends recalibration after 12 months from the date of purchase and annually thereafter.

Form TP993, Revn.3.5

Zehntner Report No.  	Page 1 of 1

Laboratory Secondary Reference Set Identification Number:

Instrument Type:

Company:

Address:

Date Tested:

Tested By:

Calibration Ref.:

Test Data:

| Panel ID: | Actual retroreflectivity (mcd/m2/lx)* |
| --- | --- |

* Uncertainty of measurement has been determined as + 7.7% of the test result, at a confidence limit of 95% with a nominal coverage factor of 2

Uncertainty of Measurement determined by and for DPTI Technical Services Group, Field Services Section.

Test procedure TP993 recommends recalibration after 12 months from the date of purchase and annually thereafter.

________________________________

Retroreflectivity Correction Table - Appendix 1 -

| Correction Values (-%) |  |  |  |  |  |  |  |  |  | RETRO (mcd) | Correction Values (+%) |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| -10 | -9 | -8 | -7 | -6 | -5 | -4 | -3 | -2 | -1 | 0 | +1 | +2 | +3 | +4 | +5 | +6 | +7 | +8 | +9 | +10 |
| 506 | 511 | 517 | 522 | 527 | 532 | 538 | 543 | 549 | 554 | 560 | 566 | 571 | 577 | 582 | 588 | 594 | 600 | 605 | 611 | 617 |
| 488 | 493 | 498 | 503 | 508 | 513 | 519 | 524 | 529 | 535 | 540 | 545 | 551 | 556 | 562 | 567 | 573 | 579 | 584 | 590 | 596 |
| 470 | 475 | 480 | 485 | 489 | 494 | 499 | 505 | 510 | 515 | 520 | 525 | 530 | 536 | 541 | 546 | 552 | 557 | 563 | 568 | 574 |
| 452 | 457 | 461 | 466 | 471 | 475 | 480 | 485 | 490 | 495 | 500 | 505 | 510 | 515 | 520 | 525 | 530 | 536 | 541 | 546 | 552 |
| 434 | 438 | 443 | 447 | 452 | 456 | 461 | 466 | 470 | 475 | 480 | 485 | 490 | 494 | 499 | 504 | 509 | 514 | 519 | 525 | 530 |
| 416 | 420 | 424 | 429 | 433 | 437 | 442 | 446 | 451 | 455 | 460 | 465 | 469 | 474 | 479 | 483 | 488 | 493 | 498 | 503 | 508 |
| 398 | 402 | 406 | 410 | 414 | 418 | 423 | 427 | 431 | 436 | 440 | 444 | 449 | 453 | 458 | 462 | 467 | 471 | 476 | 481 | 486 |
| 380 | 384 | 387 | 391 | 395 | 399 | 403 | 407 | 412 | 416 | 420 | 424 | 428 | 433 | 437 | 441 | 446 | 450 | 454 | 459 | 464 |
| 362 | 365 | 369 | 373 | 376 | 380 | 384 | 388 | 392 | 396 | 400 | 404 | 408 | 412 | 416 | 420 | 424 | 429 | 433 | 437 | 441 |
| 343 | 347 | 350 | 354 | 358 | 361 | 365 | 369 | 372 | 376 | 380 | 384 | 388 | 391 | 395 | 399 | 403 | 407 | 411 | 415 | 419 |
| 325 | 329 | 332 | 335 | 339 | 342 | 346 | 349 | 353 | 356 | 360 | 364 | 367 | 371 | 375 | 378 | 382 | 386 | 390 | 393 | 397 |
| 307 | 310 | 314 | 317 | 320 | 323 | 327 | 330 | 333 | 337 | 340 | 343 | 347 | 350 | 354 | 357 | 361 | 364 | 368 | 372 | 375 |
| 289 | 292 | 295 | 298 | 301 | 304 | 307 | 310 | 314 | 317 | 320 | 323 | 326 | 330 | 333 | 336 | 339 | 343 | 346 | 350 | 353 |
| 271 | 274 | 277 | 280 | 282 | 285 | 288 | 291 | 294 | 297 | 300 | 303 | 306 | 309 | 312 | 315 | 318 | 321 | 325 | 328 | 331 |
| 253 | 256 | 258 | 261 | 264 | 266 | 269 | 272 | 274 | 277 | 280 | 283 | 286 | 288 | 291 | 294 | 297 | 300 | 303 | 306 | 309 |
| 235 | 237 | 240 | 242 | 245 | 247 | 250 | 252 | 255 | 257 | 260 | 263 | 265 | 268 | 270 | 273 | 276 | 279 | 281 | 284 | 287 |
| 217 | 219 | 221 | 224 | 226 | 228 | 230 | 233 | 235 | 238 | 240 | 242 | 245 | 247 | 250 | 252 | 255 | 257 | 260 | 262 | 265 |
| 199 | 201 | 203 | 205 | 207 | 209 | 211 | 213 | 216 | 218 | 220 | 222 | 224 | 227 | 229 | 231 | 233 | 236 | 238 | 240 | 243 |
| 181 | 183 | 184 | 186 | 188 | 190 | 192 | 194 | 196 | 198 | 200 | 202 | 204 | 206 | 208 | 210 | 212 | 214 | 216 | 219 | 221 |
| 163 | 164 | 166 | 168 | 169 | 171 | 173 | 175 | 176 | 178 | 180 | 182 | 184 | 185 | 187 | 189 | 191 | 193 | 195 | 197 | 199 |
| 145 | 146 | 148 | 149 | 151 | 152 | 154 | 155 | 157 | 158 | 160 | 162 | 163 | 165 | 166 | 168 | 170 | 171 | 173 | 175 | 177 |
| 127 | 128 | 129 | 130 | 132 | 133 | 134 | 136 | 137 | 139 | 140 | 141 | 143 | 144 | 146 | 147 | 149 | 150 | 151 | 153 | 155 |
| 108 | 110 | 111 | 112 | 113 | 114 | 115 | 116 | 118 | 119 | 120 | 121 | 122 | 124 | 125 | 126 | 127 | 129 | 130 | 131 | 132 |
| 90 | 91 | 92 | 93 | 94 | 95 | 96 | 97 | 98 | 99 | 100 | 101 | 102 | 103 | 104 | 105 | 106 | 107 | 108 | 109 | 110 |
| 72 | 73 | 74 | 75 | 75 | 76 | 77 | 78 | 78 | 79 | 80 | 81 | 82 | 82 | 83 | 84 | 85 | 86 | 87 | 87 | 88 |
| 54 | 55 | 55 | 56 | 56 | 57 | 58 | 58 | 59 | 59 | 60 | 61 | 61 | 62 | 62 | 63 | 64 | 64 | 65 | 66 | 66 |
| 36 | 37 | 37 | 37 | 38 | 38 | 38 | 39 | 39 | 40 | 40 | 40 | 41 | 41 | 42 | 42 | 42 | 43 | 43 | 44 | 44 |
| 18 | 18 | 18 | 19 | 19 | 19 | 19 | 19 | 20 | 20 | 20 | 20 | 20 | 21 | 21 | 21 | 21 | 21 | 22 | 22 | 22 |

Form TP993A-1, Revn.3.5

Form TP993A-2 Revn.3.5

Form TP993A-3, Revn.3.5

"A"

"B"

"H"

"D"

"E"

Secondary plates calibrated for:

Identification

Number

Primary Reference

Zehntner

@ cal factor

Dedicated Transfer MX30 1621/1641

Plate ID:

Calibration Reference :

Secondary Client Reference Set

Spectra Pritchard

Actual

mcd/m2/lx

mcd/m2/lx

WORKSHEET - APPENDIX 4

Set:

Done by:

Date done:

Checked:

NOTE: CALIBRATION OF PRIMARY PLATES VALID TO NOVEMBER 2013.

Form TP993A-4 Revn.3.5

Appendix 5

Cleaning of Primary and Secondary Reference Plates

- PURPOSE

The procedure describes the processes to be used when cleaning retroreflective reference panels.

- SCOPE

This procedure applies to reference panels consisting of paint and glass beads applied over a glass substrate.

3.0	     MATERIALS AND EQUIPMENT

- Plastic container, minimum capacity 8 L

- 50mm soft synthetic bristle paint brush

- Dry compressed air supply

4.0	PROCEDURE

4.1	Place a minimum of 4 L of warm (25-35°C) fresh water in the plastic container.

4.2	Hold the panel by the edges only, wet the brush in the water and gently pass the brush from the top to the bottom of the panel, with the bristles trailing at an angle of approximately 45° to the panel.

4.3	Dip the brush into the water, and repeat the brushing action over the adjoining section of plate, repeating until all of the panel has been covered.

4.4	Rinse the panel surface under flowing fresh water, then allow to drain for a few seconds.

4.5	Dry the panel surface using compressed air.

4.6	Panel shall be tested after drying.

Note: The face of panels must never be touched by bare hands as contamination may occur.